`	``

| PREP - Amp heaters on (Heat indicator lit!) - Map out plate layout - Printout these 2 pages out separate - Label Reaction Mix tubes - 1.5mL EDLB tubes - 4 rxn per RM tube - 1 RM tubes per PGS batch (#68) - Calc amount of CLAMP MM needed - 55 µL x # batches = _________ - Use single source tube - If larger run combine MM tubes - Label strip8 tubes (or use plate) - Make 1mL 1XISS (fill out on next pg) |  | Validation PGS Form Link |
| --- | --- | --- |
|  |  | Location: MBC |
|  | Name: |  |
|  | Date and Time: |  |

RUN - Get PCR cold block out 10min before - put plate or strip8 tubes in - Get blue tube cold block out 10min before - keep all RM tubes in this - Make Reaction Mixes for all tubes: - Add 46 µL of each PGS to each RM tube (making sure to match lot#'s) - Add 55 µL of CLAMP MM to each RM Tube - Vortex for 10s and centrifuge/shake down - Add 23µL of reaction mix to each of 4 wells {at bottom, no blowout, no tip touch} - After adding RM to all reconfigure POS strips together and NEG strips together - Add NEG control (1XISS) to those columns/tubes - cap - Prep Pos Control: TPC100 - thaw 30s, add 20uL of Ultra Pure water (clean glove!) - Add POS - TPC50 to each POS tube PGS Lot# Prefix ____

|  | POS | NEG | POS | NEG | POS | NEG | POS | NEG | POS | NEG | POS | NEG |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 |
| A 1 | POS 0040 | NEG 0040 |  |  |  |  |  |  |  |  |  |  |
| B 2 | POS 0040 | NEG 0040 |  |  |  |  |  |  |  |  |  |  |
| C 3 | POS 0041 | NEG 0041 |  |  |  |  |  |  |  |  |  |  |
| D 4 | POS 0041 | NEG 0041 |  |  |  |  |  |  |  |  |  |  |
| E 5 | POS 0042 | NEG 0042 |  |  |  |  |  |  |  |  |  |  |
| F 6 | POS 0042 | NEG 0042 |  |  |  |  |  |  |  |  |  |  |
| G 7 |  |  |  |  |  |  |  |  |  |  |  |  |
| H 8 |  |  |  |  |  |  |  |  |  |  |  |  |

| PREP - Heaters on (Heat indicator lit!) - 1X Inactivation Saline Soln ready - Reaction mix strip8/plates ready - Safety Procedures: - lab coat - gloves - face mask - face shield or goggles - Alcohol wipe sample tubes |  | Amp Run Sheet Short Any System | version 1.3mod |
| --- | --- | --- | --- |

| AMPLIFICATION REACTION - Thaw Reaction Mix (if frozen) - Setup tips to align with reactions - Add 2μL from each sample tube (pipet up& down 5X, blowout in liquid, tip touch) - Heat 65°C for 25 min, set timer - Intake sample tubes in App - Put sample tubes in fridge - Remove Reactions and Let cool for 1 min before photo - Take photo in lightbox | Amp Heater: |  | Strip8 Tubes or Plates |
| --- | --- | --- | --- |
|  | Reaction Mix: Fresh |  |  |
|  | CLAMP MM Lot and EXP: |  |  |
|  | Num Reactions (including controls): |  |  |
|  | Pos Ctrl ID: TPC50 from TPC_ Ultrapure Water ID: |  |  |
|  | Neg Ctrl ID (today's 1XISS): |  |  |
|  | Initial Inconclusives?: |  |  |
| - Crop photo and apply vivid filter - Update results in App - Log Run with Form link | Notes: |  |  |